## Slide 1: Nanopore Basecalling and Read QC
GRDI-AMR2
Metagenomics Working Group
Marc-Olivier Duceppe
April 22, 2024

## Slide 2: Presentation Plan
Why Nanopore?
How to basecall the reads in “super accuracy” mode?
How to assess run / read quality?

## Slide 3: The Nanopore Technology
https://nanoporetech.com/platform/technology/basecalling
Basecalling is AI based (Deep Neural Network)
Performed from 5-mers

## Slide 4: Two Families of Sequencers
https://store.nanoporetech.com/
MinION Mk1B
Flongle
GridION
Cheaper
PromethION
P2 Solo
P2i
Higher yield and Q-score

## Slide 5: Do You Believe in Nanopore?
$1k / 87g
20bp - 100+kb
Native DNA / RNA
Real-time sequencing
Reusable flowcells
Adaptive sequencing
Growing community
Keeps changing (QA)
Error-prone
Hard to analyze
  Computationally intensive (GPU)
  Somewhat specialized tools
Pros
Cons

## Slide 6: Accuracy
Can’t quite resolve with accuracy long homopolymers stretches
Leads to lots of INDELs
Often long with more errors is better than short with less errors
AMR SNP-based:
AMR functional gene:

## Slide 7: Evolution of Basecallers
Albacore: 2017 – 2018
Guppy: 2018 – 2023
Dorado: 2023 – present
Wick RR, Judd LM, Holt KE. Performance of neural network basecalling tools for Oxford Nanopore sequencing. Genome Biol. 2019 Jun 24;20(1):129. doi: 10.1186/s13059-019-1727-y. PMID: 31234903; PMCID: PMC6591954.

## Slide 8: Read Accuracy Keeps Improving
In 2024
  R10.4.1 v14
  Dorado v0.6
  Super accuracy mode
Simplex: 99.5% (Q23)
Duplex: >99.9% (Q30)
More Q15 – Q20
https://nanoporetech.com/platform/accuracy

## Slide 9: To Guppy or not to Guppy
Why you’d still want to use Guppy
  Don’t know how to run Dorado
  You have an easier way to run Guppy
  Part of your current SOP / certified workflow
  Downstream analysis requirements
    “Old” tools with “old” data (e.g. base modification)
  Lower end GPU (< 8GB VRAM)
Why you don’t want to use Guppy anymore
  Legacy (no more updates)
  Slower
  Less accurate
  New chemistries won’t be supported

## Slide 10: Basecalling Recommendations
While sequencing: “Fast” mode
  Enables more real-time run metrics
  Do not overload host computer
Post sequencing: “Super accuracy” mode
  Higher Q-scores
  Less “fail” reads
  Better barcode assignment
  Need computer with high-end GPU
  Loose real-time benefit

## Slide 11: Guppy - Legacy
Input: Fast5 and kits (“final_summary”)
Output: Fastq and “sequencing_summary”
https://cdn.oxfordnanoportal.com/software/analysis/ont-guppy_6.5.7_linux64.tar.gz
https://github.com/duceppemo/basecall_nanopore

## Slide 12: Dorado – Two Flavours
Input: Pod5 | Fast5
Output: a single uBAM | SAM | Fastq
Needs to be demultiplexed separately
Fastq headers only contain “ReadID“
https://github.com/nanoporetech/dorado (v0.6.0)
Standalone

## Slide 13: Dorado – Two Flavours
Input: Pod5 | Fast5
Output: Fastq and “sequencing_summary.txt”
Demultiplexing built in
Fastq headers with “all” info (like Guppy)
https://cdn.oxfordnanoportal.com/software/analysis/ont-dorado-server_7.3.9_linux64.tar.gz
Basecall Server

## Slide 14: Nanopore read QC
pycoQC v2 (https://github.com/a-slide/pycoQC)
  Designed for Guppy
  Can work with Dorado
    Standalone: need to run additional commands to generate a compatible "seq_summaray.txt" file
    Server: as is

## Slide 15: Added Features
pycoQC v3 (https://github.com/duceppemo/pycoQC)
  Work in progress
  Designed to accept all dorado output file types
  Additional scripts:
    Pod5_to_seq_summary (no basecall information)
    Bam_to_seq_summary (slow)
    Fastq_to_seq_summary (recommended)
  Starting from Fastq triggers additional GC plots
  Display “All”, “Pass” or “Fail” reads

## Slide 16: Suboptimal Loading
High quality library
Not enough loaded
  Ran out after 10h
High quality library
Run stopped too early

## Slide 17: Suboptimal Sample Normalization
Uneven barcode distribution
  Poor normalization
  Adapter ligation problems
    Wrong molarity
Equal barcode distribution
  MinKNOW now uses adaptive sequencing to help balance barcodes

## Slide 18: Run Diagnosis
Very useful to find what went wrong
  DNA extraction (contaminants, short DNA fragments, etc.)
  Library prep (adapter-to-DNA ratio, sample normalization, etc.)
  Sequencing (flowcell quality, amount loaded, overheating, script error, etc.)

## Slide 19: You Can Do It!
Github repo with basic code to install and run Dorado basecall server
https://github.com/OLF-Bioinformatics/2024-04-22_GRDI_basecalling_presentation

## Slide 20: Nanopore Metagenomics Resources
https://github.com/duceppemo/mashID
  Quick species ID from metagenomics samples
  Raw reads or assemblies
  Pre-compiled DB available for bacteria
https://github.com/duceppemo/seqcounter
  Ultra-rapid preliminary AMR detection from raw reads

## Slide 21: Take-Home Message
Switch to Dorado!
Why basecall you runs using Dorado SUP?
  Access higher accuracy
  It keeps getting better (chemistry and software)
  Better downstream analysis results
DNA quality (and its size) matters
  Garbage in = garbage out
  The longer the better
  Amplified DNA sequences better (number of reads and q-score)
Bioinformaticians are not magicians
  Bad runs need to be scraped (like for diagnostics)
The truth is in the reads
  Read QC will help you find what went wrong and areas of improvement

## Slide 22: Thank You!
Q & A